## Slide 1
Integrating 12-Lead AI Reconstruction into the Sunnybrook Wearable AF Trial
Protocol Reference: Evaluation of Wearable/FitBit Monitoring to Detect Recurrent Atrial Fibrillation (v1.1)
Christopher Cheung, MD, MPH * Alex Mariakakis, PhD * Mithun Manivannan, BSc
Division of Cardiology, Sunnybrook Health Sciences Centre * Department of Computer Science, University of Toronto
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 1 / 19
Notes: Good morning Chris, good morning Alex. The purpose of today's meeting is straightforward: we already have an established clinical trial protocol—the Sunnybrook FitBit/Wearable ECG Study, authored by Chris, Alex, and our team. Our goal today is to establish how we can cleanly integrate our single-lead to 12-lead AI reconstruction research into this existing trial as a pre-specified secondary translational aim, without adding patient burden or endangering the primary clinical trial endpoints.

## Slide 2
Purpose of Today's Meeting
Protocol alignment and study integration
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 2 / 19
Clinical Trial Foundation
- Established protocol (Version 1.1, October 2025).
- Target cohort: persistent AF patients undergoing catheter ablation.
- Reference standard: 14-day continuous CardioSTAT patch at 3 and 6 months.
- Primary clinical endpoint: time to confirmed AF recurrence.
Reconstruction Constraints
- 12-lead reconstruction from single-lead input is underdetermined.
- Precordial leads exhibit amplitude and morphology errors under direct projection.
- Diagnostic reconstruction cannot serve as a primary clinical endpoint.
- Separates clinical trial validity from exploratory AI evaluation.
Translational Integration
- Routine outpatient visits already include scheduled 12-lead ECGs.
- Synchronous smartwatch and 12-lead acquisition adds <= 5 minutes.
- Evaluates reconstruction and biomarker recovery as a secondary aim.
- Generates a prospectively paired clinical and wearable dataset.
Objective: Align on the secondary data collection protocol and freeze translational evaluation metrics.
Notes: Here is the layout for today. Box 1 is our foundation: Chris's clinical protocol is locked. It evaluates smartwatch monitoring versus continuous CardioSTAT patch for post-ablation AF recurrence. Box 2 is our methodological reality check: we know 12-lead reconstruction does not work with uniform diagnostic precision across all leads and patients; staking the trial on reconstruction would be scientifically unsound. Box 3 is our opportunity: during routine clinical visits, we can capture simultaneous paired watch and 12-lead recordings. This establishes a publishable secondary translational AI aim while protecting the primary trial.

## Slide 3
The Established Clinical Protocol (v1.1)
Protocol v1.1 design and endpoints
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 3 / 19
Trial Design & Patient Population
- Single-center prospective cohort at Sunnybrook Arrhythmia Clinic.
- Patients aged >= 18 with persistent AF undergoing catheter ablation or cardioversion.
- Smartwatch monitoring: daily 30-second ECG and automated rhythm notifications for 6 months.
- Reference standard: 14-day continuous CardioSTAT patch at 3 and 6 months post-procedure.
Endpoints & Sample Size
- Primary outcome: time to recurrent AF detection (smartwatch vs. patch).
- Secondary clinical outcomes: heart rate trends, heart rate variability, and cumulative AF burden.
- Sample size: N = 96 patients (80% power, alpha = 0.05, 10% non-inferiority margin).
- Vanguard phase: initial 20 consecutive patients to assess protocol adherence and data pipeline.
Primary trial endpoints and clinical workflows remain unchanged under this proposal.
Notes: Chris designed this study to evaluate whether daily smartwatch ECG recordings and photoplethysmography irregular rhythm notifications detect recurrent AF earlier than standard-of-care intermittent patch monitoring. Patients receive a 14-day CardioSTAT patch at 3 months and 6 months post-procedure. The sample size is 96 patients, starting with a 20-patient vanguard. This clinical core remains completely sovereign.

## Slide 4
Why Reconstruction Cannot Be the Primary Endpoint
Methodological constraints of single-lead reconstruction
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 4 / 19
Evaluation Dimension | Primary Clinical Trial Standard | Single-Lead 12-Lead Reconstruction
Mathematical Problem | Time-to-event survival analysis | Underdetermined inverse problem (1 -> 12 projection)
Empirical Performance | Wearable AF sensitivity ~ 94% (Abu-Alrub 2022) | Precordial leads (V2-V4) exhibit variance loss (r ~ 0.74, MAE ~ 0.13 mV)
Reported Literature | Guideline-endorsed recurrence metrics (2024 EHRA/HRS) | Presacan et al. (2025): amplitude compression and regression to mean
Hardware Domain Gap | Verified against 14-day CardioSTAT patch | Hu et al. (2026): hospital limb leads differ from dry-electrode wrist signals
Trial Execution Risk | Low: clinical utility evaluable independently | High: staking trial success on 12-lead fidelity introduces regulatory risk
Methodological Summary: Single-lead reconstruction does not replace diagnostic 12-lead acquisition. Framing reconstruction as a primary endpoint introduces substantial clinical and regulatory risk, whereas a secondary translational aim leverages paired data without impacting trial governance.
Notes: Reviewers and clinical colleagues will rightly push back if anyone claims a smartwatch can replace a diagnostic 12-lead machine today. Electrophysiologically, reconstructing 12 leads from one lead is an underdetermined inverse problem with known variance compression. If a clinical trial staked its primary regulatory endpoint on 12-lead reconstruction accuracy, it would introduce severe risk. However, that does not mean the data is useless—it means reconstruction belongs strictly as a secondary translational aim.

## Slide 5
The Clean Integration: A Rigorous Two-Track Protocol
Separation of primary clinical and secondary AI aims
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 5 / 19
Track 1: Primary Clinical Evaluation (Aim 1)
- Evaluates whether smartwatch monitoring improves time to detection of recurrent AF.
- Reference standard: 14-day CardioSTAT patch at 3 and 6 months (iCentia analysis).
- Sized for AF recurrence detection sensitivity and non-inferiority (N = 96).
- Inviolable primary endpoint governing trial reporting.
Track 2: Secondary Translational AI (Aims 2 & 3)
- Aim 2: Benchmark 12-lead waveform recovery from wrist signals using paired clinic data.
- Aim 3: Evaluate within-patient repolarization tracking (delta QTc) using baseline priors.
- Secondary, exploratory endpoints with no effect on primary trial conclusions.
- Independent computational and methodological analyses.
Preserves primary clinical endpoints while establishing a prospectively paired validation dataset.
Notes: This architecture cleanly separates the two efforts. Track 1 is Chris's primary clinical trial: time to recurrent AF detection, backed by the 14-day CardioSTAT patch. This guarantees the study will produce high-impact clinical papers regardless of AI performance. Track 2 is Alex's and our translational AI research: we pre-specify 12-lead reconstruction as a secondary translational aim. We use the in-clinic paired recordings to benchmark transfer functions, domain adaptation, and patient-specific priors. This protects trial integrity while maximizing scientific return.

## Slide 6
Multitask Auxiliary Supervision (Ansari et al., 2026)
Spatial representation learning for single-lead models
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 6 / 19
Stanford 3DRECON-QT Framework
- Reference: Ansari, Rogers et al., Circulation (January 2026).
- Multitask objective: simultaneously reconstruct 12 leads and estimate clinical biomarkers.
- Auxiliary 12-lead loss enforces 3D cardiac vector representations in the encoder.
- Removing spatial supervision reduced downstream biomarker correlation from r = 0.73 to 0.04.
Application to Single-Lead AF Surveillance
- Daily single-lead smartwatch ECG monitors for AF recurrence.
- Lead I captures frontal plane potentials, with limited sensitivity to orthogonal atrial vectors.
- Auxiliary 12-lead reconstruction head encourages representation of full 3D vector loops.
- Supports atrial feature extraction without altering primary single-lead inference inputs.
Auxiliary 12-lead reconstruction acts as an inductive regularizer during pretraining.
Notes: This slide connects our project to recent literature published in Circulation by Albert Rogers' group at Stanford (Ansari et al., 2026). Ansari developed 3DRECON-QT, showing that training an auxiliary 12-lead reconstruction head forces the single-lead encoder to learn 3D cardiac vector loops. In their ablation study, removing the 12-lead spatial conditioning collapsed downstream clinical biomarker prediction from r=0.73 down to 0.04. We apply this paradigm to our trial: at home, patients record daily single-lead smartwatch ECGs for our primary endpoint: time to AF recurrence. Lead I alone often misses vector components oriented perpendicular to the frontal plane. Training an auxiliary 12-lead reconstruction head helps the encoder preserve 3D atrial vector dynamics.

## Slide 7
Figure 1: The ECG-AIM Model Architecture
Network architecture and feature extraction
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 7 / 19
Architecture Components
- 50 ms temporal patches project single-lead input (1x5000) into token embeddings.
- Continuous Morlet wavelet bank across 32 frequency scales (0.5-45 Hz).
- Combined TimeSformer spatiotemporal attention and 1D convolutions.
- Sigmoid-gated residual fusion: Z = Z_t + sigma(W[Z_t, Z_w]) * Z_w.
- 8-layer transformer encoder (d=768, 12 attention heads).
- Axial transformer decoder resolving temporal and cross-lead spatial features.
Notes: This diagram shows the architecture of our champion model: conv15e_A0_wave_noSSL_gated_add. Notice the core components: 1. Patch tokenizer slicing the Lead I input into 50 ms tokens. 2. A continuous Morlet wavelet bank across 32 frequency scales from 0.5 to 45 Hz. 3. A hybrid TimeSformer and 1D-CNN backbone. 4. Our gated-add fusion module that dynamically modulates wavelet energy via a sigmoid gate. 5. An 8-layer transformer encoder. 6. An axial transformer decoder producing both continuous 12-lead waveforms and explicit fiducial delineations.

## Slide 8
Figure 2: ECG-AIM Multi-Task Training Pipeline
Single-lead simulation and objective functions
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 8 / 19
Notes: This figure illustrates how we trained the ECG-AIM model before testing on single-lead signals. First, on the left, we ingest full 12-lead hospital records from PTB-XL and MIMIC-IV. We simulate single-lead input by masking out the other 11 leads. Second, the single-lead signal passes through our dual-branch wavelet transformer backbone to predict continuous 12-lead waveforms and P-QRS-T boundary delineations. Third, looking at the loss formulation: we supervise waveform reconstruction with a composite loss (MSE, Pearson correlation, Wyatt ST-T phase loss, and spectral loss) plus Einthoven limb consistency, alongside multi-task cross-entropy and Dice loss on delineation. Importantly, during training, the single lead is a synthetic proxy derived from hospital ECGs on supine patients. It lacks dry-electrode contact noise and wearable filtering. This gap motivates the need for our Sunnybrook paired trial data to calibrate and validate the model on clinical hardware.

## Slide 9
Empirical Validation of ECG-AIM: Benchmark Performance
Comparative benchmark against 3DRECON-QT
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 9 / 19
Clinical Evaluation Dimension | Factorial Baseline | Stanford 3DRECON-QT | ECG-AIM (Ours) | Statistical Contrast
Precordial Chest Mean Pearson r | 0.570 | 0.732 | 0.776 | Delta = +0.206 (p < 10^-300), Ranked #3 of 55
Anterior Lead V3 Pearson r | 0.382 | 0.670 | 0.739 | Delta = +0.357 (p < 10^-300), Anterior wall view
QRS Duration Error (MAE / Median) | 13.85 ms | 10.26 ms | 10.05 / 8.0 ms | Depolarization tracking
Conduction Delay Concordance (>120ms) | 90.7% | 94.3% | 94.3% | aOR = 1.84 (p < 10^-17), Specificity: 96.5%
Sokolow-Lyon LVH Voltage MAE | 1.13 mV | 0.63 mV | 0.72 mV | ICC = 0.714, Hypertrophy quantification
EchoNext Structural Heart Disease | 68.8% | 46.5% | 71.6% | Preserves chamber echocardiographic pathology
Time-Frequency Feature Preservation
Continuous Morlet wavelets preserve high-frequency QRS details and ST-T repolarization without baseline distortion, improving chest lead correlation over static CNN baselines (r = 0.776 vs. 0.732).
Generalization to Structural Phenotypes
On the EchoNext cohort, 3DRECON-QT demonstrated reduced structural concordance (46.5% overall, 33.3% on reduced LVEF). The gated wavelet transformer maintains 71.6% concordance across echocardiographic phenotypes.
Notes: This slide presents the benchmark across models evaluated with patient-clustered MMRM and GEE. Compared to the Stanford 3DRECON-QT model, our model achieves higher precordial chest lead correlation—0.776 vs 0.732—and a Lead V3 correlation of 0.739 vs 0.670. On echocardiographic structural heart disease in EchoNext, Ansari's model dropped to 46.5% concordance because their scalar QT loss did not constrain chamber geometry. Our wavelet-gated transformer maintains 71.6% structural concordance and tracks QRS duration with a median error of 8.0 ms.

## Slide 10
Zero-Shot Multi-Cohort Generalization & Flexible Lead Inputs
Evaluation across variable lead configurations and external cohorts
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 10 / 19
Dynamic Lead Masking (M_obs)
- Stochastic lead masking during training enables arbitrary lead subsets at inference without retraining.
- Ambulatory screening (Lead I): PTB-XL AF AUROC: 0.952; Macro 150-task AUROC: 0.821.
- Clinical evaluation (Leads I, II, V2): Orthogonal vertical and horizontal leads resolve dipole ambiguity.
- Missing lead correlation increases to r = 0.892 - 0.945; QRS duration MAE reduced to 8.12 ms.
Cohort & Clinical Metric | Stanford 3DRECON | 1-Lead ECG-AIM | 3-Lead ECG-AIM
PTB-XL Foundation (N = 21,799) |  |  | 
ECGFounder 150-Task AUROC | 0.7720 | 0.8208 | 0.8409
AF Recurrence Detection AUROC | 0.9140 | 0.9520 | 0.9812
Ventricular Tachycardia AUROC | 0.8840 | 0.9350 | 0.9894
Conduction Delay (>120ms) Concord. | 94.3% | 94.6% | 96.8%
QRS Duration MAE | 10.26 ms | 10.05 ms | 8.12 ms
EchoNext Structural (N = 100,000) |  |  | 
12 Deep Phenotypes Macro AUROC | 0.5890 | 0.7374 | 0.7542
Reduced LVEF (<=45%) Concord. | 33.3% | 77.7% | 81.4%
LV Hypertrophy (>=13mm) Concord. | 30.2% | 68.9% | 73.5%
Sunnybrook 10-Wire XMLs (N = 20) |  |  | 
Physical Limb Leads | Uncalibrated | Preserved | r = 0.938
Biological Noise Floor (~18 uV) | Attenuated | Preserved | Preserved
A single model adapts to available inputs: single-lead ambulatory screening or 3-lead clinical triage.
Notes: This slide highlights model evaluation under flexible lead inputs. Because we train with dynamic lead token masking, the exact same model weights accept variable combinations of leads. At home, single-lead Lead I achieves an ECGFounder Macro AUROC of 0.821 and detects AF with an AUROC of 0.952. When a patient attends clinic, adding Leads II and V2 resolves orthogonal planes without changing model weights: AF detection AUROC reaches 0.981, VT detection reaches 0.989, and QRS duration error drops to 8.12 ms. On EchoNext, the model maintains 75.4% AUROC and 81.4% LVEF concordance.

## Slide 11
Single-Lead Reconstruction: Smartwatch (Lead I) vs. Patch (Lead II)
Reconstruction of PTB-XL Record #10283 from single-channel inputs
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 11 / 19
Lead I Input (Transverse Axis)
- Single-channel input: Lead I (wrist-to-wrist vector).
- 11 reconstructed leads yield mean r = 0.941.
- Precordial correlation: V1 (r=0.995), V2 (r=0.989), V3 (r=0.988).
Lead II Input (Vertical Axis)
- Single-channel input: Lead II (vertical torso vector).
- 11 reconstructed leads yield mean r = 0.919.
- Inferior leads: aVF (r=0.992), III (r=0.963), V2 (r=0.966).
Lead I input provides strong anterior precordial recovery (r > 0.98), whereas Lead II prioritizes inferior leads.
Notes: Here is the single-lead comparison between sensor modalities on unseen PTB-XL Record 10283. On the left, we input single-lead Lead I—the wrist-to-wrist vector acquired by smartwatch devices. The model infers the other 11 leads with a record mean correlation of 0.941. Notice the reconstruction of precordial chest leads: V1 is 0.995, V2 is 0.989, and V3 is 0.988. On the right, we input single-lead Lead II—the vertical vector typically captured by adhesive chest patches. Lead II achieves a record mean correlation of 0.919, with recovery of inferior leads like aVF at 0.992 and III at 0.963.

## Slide 12
Impact of Input Dimension: 1-Lead vs. 3-Lead Input
Reconstruction of PTB-XL Record #10283 under single-lead and 3-lead configurations
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 12 / 19
Single-Lead Input (Lead I)
- Single-channel input: Lead I.
- 11 reconstructed leads yield mean r = 0.941.
- Designed for continuous ambulatory surveillance.
Three-Lead Input (I, II, V2)
- Three-channel input: orthogonal triad (I, II, V2).
- 9 reconstructed leads yield r = 0.980 - 1.000.
- Limb completion: r = 1.000; QRS MAE = 8.12 ms.
Adding two orthogonal leads (II and V2) constrains both planes, improving precordial and limb fidelity.
Notes: This comparison shows scaling from single-lead to 3-lead input. On the left is single-lead Lead I input with a mean correlation of 0.941. While suitable for outpatient rhythm surveillance, a single lead is mathematically limited in resolving orthogonal planes simultaneously. On the right, adding Lead II and V2 provides orthogonal axes. Einthoven completion yields 1.000 across all limb leads, and precordial chest leads reach correlations between 0.980 and 0.999, reducing QRS duration error to 8.12 ms.

## Slide 13
External Cohort Validation: Sunnybrook In-Clinic XMLs
Zero-shot evaluation on Record ECG004 under 1-lead and 3-lead configurations
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 13 / 19
Single-Lead Input (Lead I)
- Single-channel input: Einthoven Lead I.
- 11 reconstructed leads yield mean r = 0.865.
- Selected leads: aVR (r=0.966), V5 (r=0.934), V6 (r=0.929), II (r=0.925).
Three-Lead Input (I, II, V2)
- Three-channel input: orthogonal triad (Leads I, II, V2).
- 9 reconstructed leads yield mean r = 0.938.
- Limb completion: r = 1.000; precordials: V1 (r=0.973), V6 (r=0.957).
Zero-shot evaluation on in-clinic XMLs preserves physiological amplitude (> 2.1 mV in V2) without recalibration.
Notes: This slide shows zero-shot evaluation on Sunnybrook hospital XML recordings. Record ECG004 displays clinical voltages—over 1.4 mV in Lead II and 2.1 mV in V2. On the left, using single-lead Lead I across physical hospital hardware with un-derived limb leads and natural noise, the model achieves a record mean correlation of 0.865 (Lead II at 0.925, aVR at 0.966, V5 at 0.934). On the right, with the 3-lead triad, mean correlation increases to 0.938. Limb leads reach 1.000, V1 reaches 0.973, and V6 reaches 0.957.

## Slide 14
External Cohort Validation: EchoNext Structural Cohort
Zero-shot evaluation on Record #10 under 1-lead and 3-lead configurations
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 14 / 19
Single-Lead Input (Lead I)
- Single-channel input: Lead I.
- 11 reconstructed leads yield mean r = 0.809.
- Selected leads: aVR (r=0.967), aVL (r=0.946), V1 (r=0.942), V5 (r=0.915).
Three-Lead Input (I, II, V2)
- Three-channel input: orthogonal triad (Leads I, II, V2).
- 9 reconstructed leads yield mean r = 0.964.
- Limb completion: r = 0.986 - 1.000; precordials: V1 (r=0.987), V6 (r=0.965).
Preserves ventricular waveforms on echo-paired tracings (Record #10: peak voltages 0.9 - 1.4 mV; LVEF concordance 81.4%).
Notes: This slide shows external validation on the Stanford EchoNext cohort of paired echocardiogram-ECG patient tracings. Examining Record 10, featuring normal voltages up to 1.40 mV in V4 and 1.15 mV in V2: On the left, 1-lead reconstruction achieves a record mean correlation of 0.809, with aVR at 0.967 and chest leads V1 and V5 at 0.942 and 0.915. On the right, adding the 3-lead triad increases mean correlation to 0.964, with limb tracking between 0.986 and 1.000, and precordial tracking reaching 0.987 on V1 and 0.965 on V6. Preserving ST-T morphology and QRS amplitude allows downstream models to achieve 81.4% concordance on reduced ejection fraction.

## Slide 15
Data Collection Methodology: Dual-Stream Protocol
Ambulatory monitoring and in-clinic paired acquisition
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 15 / 19
Stream A: Ambulatory Surveillance
- Daily 30-second smartwatch ECG (Lead I) for 6 months post-ablation.
- 14-day continuous CardioSTAT patch worn at 3 and 6 months.
- Primary trial endpoints: time to AF recurrence, AF burden, and symptom correlation.
- Automated data transmission to Sunnybrook secure study server.
Stream B: In-Clinic Paired Protocol
- Four scheduled clinical visits: baseline, post-procedure, 3 months, and 6 months.
- Simultaneous recording of 12-lead ECG and smartwatch Lead I (delta t = 0).
- Standardized acquisition adds <= 5 minutes to routine visit.
- Requires no additional clinic visits or invasive procedures.
Stream A fulfills primary trial objectives; Stream B provides synchronous paired data for secondary analysis.
Notes: This slide outlines the data collection workflow. It operates as two synchronized streams: Stream A is the ambulatory home surveillance from Protocol v1.1. Patients wear the watch daily for 6 months and wear the 14-day continuous CardioSTAT patch at 3 and 6 months. This feeds directly into our primary clinical efficacy endpoint. Stream B is the in-clinic paired protocol. Patients are already at Sunnybrook for four standard clinical visits: baseline pre-ablation, procedure day, 3 months, and 6 months. During the routine clinical 12-lead ECG, the research coordinator records a 30-second smartwatch tracing simultaneously. It adds less than 5 minutes to routine care and creates synchronized ground-truth pairs.

## Slide 16
Paired In-Clinic Dataset Design
Comparison with prior retrospective evaluations
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 16 / 19
Methodological Dimension | Ansari et al. (2026) | Sunnybrook Prospective Protocol
Evaluation Cohort | Retrospective CareLink archive (N = 26) | Prospective cohort (N = 96 patients, ~384 paired visits)
Temporal Offset | +/- 7 days between device and 12-lead ECG | Synchronous acquisition (delta t = 0, simultaneous)
Physiological State | Uncontrolled variations in hydration and medication | Identical resting state (supine, matched heart rate)
Rhythm States | Uncontrolled outpatient recordings | Defined clinical milestones (baseline AF, post-procedure, follow-up)
Methodological Role | Device-specific feasibility check | Prospective benchmark for domain adaptation and prior models
Significance of Prospective Paired Data: Most published ECG reconstruction studies evaluate models on synthetically masked 12-lead recordings. Synthetic masking does not account for dry-electrode interfaces, skin impedance, or wearable filtering. Synchronous in-clinic acquisition provides an empirical standard for evaluating wearable reconstruction models.
Notes: This slide contrasts our prospective paired acquisition with the retrospective analysis in the Stanford paper. Ansari and Rogers searched through Medtronic CareLink transmissions, identifying 26 patients who had a clinical 12-lead within plus-or-minus 7 days. Over a week, autonomic tone, electrolytes, and medications can vary. In our trial, we acquire simultaneous pairs across 96 patients at 4 defined clinical visits. Furthermore, this cohort captures longitudinal transitions: pre-ablation AF, acute post-procedure rhythm, and 3-to-6-month follow-up. This directly addresses the gap between synthetic lead masking and physical wearable recording.

## Slide 17
Figure 3: Integrated Clinical-AI Trial Pipeline
Trial workflow and computational integration
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 17 / 19
Integration Stages
- Pretraining on public cohorts (PTB-XL, MIMIC-IV, Icentia11k) to establish baseline representations.
- Data acquisition across three streams: smartwatch daily rhythm and in-clinic paired Lead I; 14-day continuous CardioSTAT patch; hospital 12-lead machine.
- Hardware transfer function estimation and adapter calibration on paired recordings.
- Evaluation on distinct endpoints: primary AF recurrence and burden vs. CardioSTAT patch; secondary 12-lead reconstruction and repolarization tracking.
Notes: Figure 3 shows how the clinical trial and AI engineering integrate. Stage 1 represents foundation models pretrained on public datasets. Stage 2 represents data acquisition: smartwatch, 14-day continuous CardioSTAT patch, and hospital 12-lead machine. Stage 3 uses the paired in-clinic data to calibrate the hardware transfer function and adapt the model. Stage 4 shows the separation of endpoints: primary clinical AF surveillance vs. secondary translational 12-lead reconstruction.

## Slide 18
Sample Size, Statistical Power & Clustered Modeling
Sample size justification and statistical analysis plan
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 18 / 19
Primary Clinical Power (Protocol v1.1)
- Primary outcome: time to confirmed AF recurrence.
- Hypothesis: smartwatch surveillance is non-inferior to 14-day patch monitoring.
- Event rate assumptions: 20% recurrence with patch, 25% with smartwatch.
- Target power: 80% power, alpha = 0.05, 10% non-inferiority margin (N = 96 patients).
- Vanguard phase: initial 20 consecutive patients to evaluate workflow feasibility.
Secondary AI Precision & Analysis Plan
- Sample yield: 96 patients across 4 clinic visits provides >2,500 paired complexes.
- Precision target: bounds QTc estimation error within 95% CI margin of +/- 3.5 ms.
- Within-patient beat clustering violates independent observation assumptions.
- Linear mixed-effects models (unstructured covariance) and generalized estimating equations (GEE).
Primary sample size is powered on clinical event rates; secondary analysis accounts for repeated measures.
Notes: This slide outlines sample size and statistical considerations. The primary clinical trial sample size of 96 patients is calculated directly from expected AF recurrence rates—20% on patch vs 25% on smartwatch with an 80% power target. For the secondary AI aim, 96 patients across multiple clinic visits provides over 2,500 paired cardiac complexes. Per Riley's 2024 BMJ guidelines, this is sufficient to bound QTc measurement error within a 3.5 ms margin of error. We specify linear mixed-effects models and generalized estimating equations to adjust for intra-patient beat clustering.

## Slide 19
Action Items & Alignment for Today's Meeting
Decisions for protocol finalization
Sunnybrook Arrhythmia Trial Protocol * 12-Lead AI Integration 19 / 19
Decisions for Today's Meeting
1. Protocol structure: Approve formal separation of Aim 1 (clinical trial) and Aim 2 (translational AI).
2. Paired recording protocol: Approve 5-minute simultaneous in-clinic acquisition procedure.
3. Secondary biomarker selection: Finalize priority clinical biomarker (QTc monitoring vs. conduction intervals).
4. Vanguard enrollment: Authorize start of the 20-patient feasibility vanguard.
Expected Outcomes
- Clinical outcomes: Prospective evaluation of smartwatch surveillance for post-ablation AF recurrence.
- Methodological outcomes: Benchmark dataset and algorithms for wearable 12-lead reconstruction.
- Trial integrity: Clear division between primary regulatory endpoints and exploratory AI objectives.
Next Step: Finalize protocol amendment text and initiate vanguard enrollment.
Notes: To close our meeting, here are the 4 decisions we need to agree on: First, confirm the two-track protocol amendment: Aim 1 primary AF recurrence, Aim 2 secondary 12-lead reconstruction. Second, approve the 5-minute simultaneous recording procedure during in-clinic visits. Third, decide which secondary clinical biomarker Chris wants to prioritize for the paired analysis—QTc interval tracking is our recommendation. Fourth, approve the 20-patient vanguard cohort. Thank you Chris and Alex; let us open the floor for discussion.